## Slide 1: Image reconstruction 28/7/2024, TPPT project
ICNAS
Nuno C. Ferreira

## Slide 2: Summary
NEMA-like FWHM and FWTM for linear source in specific points of the FOV (9 positions):
(Note: the FBP reconstructed images were rotated 180 degrees to make it correspond to the positions in this diagram.)

## Slide 3: Limitations
Limitation:
  Non-TOF only
    should not be very different with TOF
Differences relative to the NEMA standard
  Plot range should be approx 2x the FWHM
    this is according to the NEMA standard, but I disregarded that. I am plotting a fixed range (20mm on each side of the peak).
  Positions and type of source is not as specified in NEMA
    here I am using the 9 positions of the previous slide.

## Slide 4
This NEMA requirement has now been implemented:
Parabolic fit to find the peak based on the max and one neighbor on each side
For reference, if needed, the position of the peak that was found using the parabolic fit is shown as a very subtle orange point in the plots. Sometimes it is indistinguishable from the maximum of the plot.
these undershoots in the Y profiles are mainly due to the incomplete ring of TPPT, which is inconsistent with what the FBP algorithm expects (full-ring geometry)

## Slide 5: Method
Data was reconstructed with FBP (no apodization or smoothing)
  listmode data was binned into non-oblique sinograms (no rebinning was used). These sinograms were added, resulting in one non-oblique sinogram.
The Y profiles across the peak have negative values
  they are due to the incomplete ring geometry (the FBP algorithm expects data from a full ring geometry)
Example of FBP reconstruction
Detector head
Detector head

## Slide 6: Regarding the sinograms:
Due to the limited-angle geometry, they are incomplete, i.e. not all angular positions have data. Example:
(a line source perpendicular to the image plane would generate a sinosoidal line in the sinogram, but due to the TPPT limited-angle geometry we only see part of it.)

## Slide 7: Regarding the FBP algorithm:
No zero-padding was applied, to conform to the NEMA standard.
  in yesterdays’ slides, zero padding was being applied in my FBP implementation.

## Slide 8: Results
Summary of the FWHM and FWTM obtained with different image sizes (FBP images were reconstructed with: 200x200, 300x300, 400x400 and 500x500 pixels)

## Slide 9: Summary
All the best FWHM and FWTH were obtained with the 500x500 images.
  Best FWHM_y: 0.71 mm
  Best FWHM_x: 1.61 mm

## Slide 10
Number of pixels in reconstructed image (NxN)
due to noise in the image

## Slide 11: Results
200x200 pixels FBP reconstructed images
X and Y FWHM and FWTM
Worse spatial resolutions are expected, but less image noise

## Slide 21: Results
500x500 pixels FBP reconstructed images
X and Y FWHM and FWTM
Better spatial resolutions are expected, but more image noise

## Slide 27
Best FWHM_y

## Slide 30
Best FWHM_x

## Slide 31: Results
200x200 pixels
sinograms and reconstructed images (FPB)
Other image sizes (300x300, 400x400, 500x500) are available in a .zip file.

## Slide 32
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 33
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 34
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 35
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 36
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 37
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 38
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 39
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 40
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 41: Results
500x500 pixels
sinograms and reconstructed images (FPB)
Other image sizes (300x300, 400x400, 500x500) are available in a .zip file.

## Slide 42
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 43
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 44
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 45
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 46
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 47
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 48
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 49
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 50
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 51: Results
Effect of the number of pixels on the sinograms and reconstructed images (FPB):
one example.

## Slide 52
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 53
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 54
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 55
the reconstructed image was rotated 180 degrees to correspond to the diagram 

## Slide 56: Discussion
To report the NEMA spatial resolution, we should have a point source in the positions shown in the figure. We have positions at center FOV and at 3 cm from the center FOV, along X and Y (it is not very different).
Our sources are different and the positions were also different. So, we are not yet reporting NEMA spatial resolutions, but figures that indicate what we can expect in close to NEMA conditions.
I have summed the data corresponding to all the planes (line source; NEMA requires point source). So, I did not calculate the spatial resolution for different positions axially (NEMA would give two values, at center and at ¼ axial FOV). I can give those different axial values, if needed, but the conditions are not NEMA (line source instead of point source).
1 cm vertically from the center